- PF400 Robot Plug-In for TCS

- Version 2.0B1, July 2012

# Overview

This document describes the PF400 Robot Plug-In for the Precise Automation TCP Command Server (TCS) application.  It contains a number of commands specific to the PF400 robot and its electronic gripper.

To use this plug-in,  load the Tcp_cmd_server_pf400 project which consists of the standard TCS plus the plug-in file PF400.gpl

Within this file, the module PF400 contains commands for the PF400 robot, and the module SSGRIP contains commands for the Precise Spring Servo Gripper.

# PF400 Commands

The table below describes the commands contained in this plug-in.

See the TCP Command Server User's Guide for information about the standard TCS commands.

These commands apply to a PF400 Robot with an optional Precise Servo Gripper and an optional linear traverse rail.

| Command Case ignored | Arguments | Description, responses are for PC mode. Telnet mode is similar but more descriptive. |
| --- | --- | --- |
| changeconfig | arg1: Grip mode. 0 = do not change gripper 1 = open gripper 2 = close gripper | Change Robot configuration from Righty to Lefty or vice versa. Does not include checks for collision inside work volume of the robot. Can be customized by user for their work cell configuration |
| freemode | arg1: Axis selection. -1 = Unfree all axes. 0 = Free all axes. > 0 = Free just this axis. | Activates or deactivates free mode. |
| GraspData | arg1-arg3 must all be omitted or all specified. arg1: The plate width in mm. arg2: The finger speed during grasp where 100 means 100%. arg3: The gripper squeezing force, in Newtons. | Gets or sets the data to be used for the next force-controlled PickPlate command grip operation. If all arguments are omitted, returns the response is: plate width, finger speed %, grasp force. This data remains in effect until the next GraspData command or the system is restarted. |
| gripclosepos | arg1: If omitted, the gripper close position is returned. If specified, contains the new gripper close position. | Gets or sets the gripper close position for PF400 servo gripper. The close position may be changed by a force-controlled grip operation. |
| gripopenpos | arg1: If omitted, the gripper open position is returned. If specified, contains the new gripper open position. | Gets or sets the gripper open position for PF400 servo gripper. |
| gripper | arg1: Grip mode. 1 = Open gripper 2 = Close gripper | Opens or closes the servoed gripper. |
| MoveRail | arg1: The destination station ID. Only used if arg3 is omitted. arg2: Optional mode. 0 or omitted = cancel any pending MoveRail. 1 = Move rail immediately. 2 = Move rail during next pick or place. arg3: Optional. If specified, use this value as the rail destination rather than the station location. | Moves the optional linear rail. The rail may be moved immediately or simultaneously with the next pick or place motion. The location may be associated with the station or specified explicitly. |
| movetosafe | none | Moves PF400 Robot to Safe Position. Does not include checks for collision 3rd party inside obstacles inside the work volume of the robot |
| pickplate | arg1: Station ID, from 1 to 100. Corresponds to TCS locidx value + 1 | Moves to a predefined position and picks up plate. If the arm must change configuration, it automatically goes through the Park position. At the conclusion of this routine, the arm is left gripping the plate and moving to the nest approach position. Use Teach function to teach station pick point. |
| placeplate | arg1: Station ID from 1 to 100. Corresponds to TCS locidx value + 1 | Moves to predefined position and place a plate. If the arm must change configuration, it automatically goes through the Park position. At the conclusion of this routine, the arm is left gripping the plate and moving to the nest approach position. Use Teach function to teach station place point. Returns -1 if the plate was successfully grasped or force control was not used. Returns 0 if the force-controlled gripper detected no plate present. |
| Rail | arg1: Station ID from 1 to 100. Corresponds to TCS locidx value + 1 arg2: Optional. If omitted, the rail value is returned. If specified, this is the new rail position. | Gets or Sets the position of the optional rail axis that is associated with a station. The station rail data is loaded and saved by the LoadFile and UpdateFile commands. |
| teachplate | arg1: Station ID, from 1 to 100. Corresponds to TCS locidx value + 1 arg2: Optional. The Z Clearance value. If omitted, a value of 50 is use. If specified and non-zero, this value is used. | Sets the plate location to the current robot position and configuration. The location is saved as Cartesian. Z clearance must be high enough to withdraw the gripper. |
| StationType | arg1: If omitted, the station type is returned. If specified, contains the new station type. 0 = horizontal, 1 = vertical. | Gets or sets the station type for PF400 robot for the pickplate and placeplate commands. |

# SSGRIP Commands

These commands operate the Precise Spring Servo Gripper for use in lab automation applications.

| Command Case ignored | Arguments | Description, responses are for PC mode. Telnet mode is similar but more descriptive. |
| --- | --- | --- |
| HomeAll_IfNoPlate | none | Tests if the gripper is holding a plate. If not enable robot power and home all robots. Returns -1 if no plate detected and the command succeeded. Returns 0 if a plate was detected. |
| GraspPlate | arg1: Plate width, in mm. Should be accurate to within about 1 mm. arg2: Percent speed to close fingers. 1 to 100. arg3: Maximum gripper squeeze force, in Nt. | Grasps a plate with limited force. The actual closing distance is about 1 mm smaller than arg1 so that the PID servos generate a gripping force. Due to noise, friction, and the spring, actual gripping force may vary somewhat. Returns -1 if the plate has been grasped. Returns 0 if the final gripping force indicates no plate. |
| ReleasePlate | arg1: Open width, in mm. arg2: Percent speed to open fingers. 1 to 100. arg2: Optional. The standard InRange profile property for the gripper open move. If omitted, a zero value is assumed. | Releases the plate after a GraspPlate command. Opens the gripper to the specified width and cancels the force limit once the fingers are open. |
| IsFullyClosed | none | Tests if the gripper is fully closed by checking the end-of-travel sensor. Returns -1 if the gripper is within 2mm of fully closed. Otherwise returns 0. |

# Customization

## Customization variables

The following table summarizes the variables in this plug-in that you might want to modify.

| Variable | Description |
| --- | --- |
| USE_FORCE_GRIPPER | Set to True to control the gripper force during PickPlate and PlacePlate. The plate will be handled by internally using GraspPlate and ReleasePlate Set to False to control the gripper position with no special force considerations. |
| Z_ABOVE | The height in mm above the nest to which the robot moves when approaching or departing from the nest. The standard value is 5 mm. |
| Z_CLEARANCE | The default distance in mm to withdraw the gripper when approaching or departing from the nest. The standard value is 50 mm. |
| RAIL_TOLERANCE | The tolerance to be used to determine if the optional linear rail is at the correct station. The standard value is 10 mm. |

## Custom hook routines

The following routines are called by TCS to perform custom operations.

### Hook_InitCommand

### This routine is called when a new command object is created for a robot.

### Hook_LoadFile

### This routine is called when the LoadFile command loads data from a .gpo file.  It reads the rail positions from variables named tcs_rail_robotnum_index

where:

robotnum is the number of the robot being used.

index is the station number - 1.

### Hook_UpdateFile

### This routine is called when the UpdateFile command stores data to a .gpo file.  It writes the rail positions to variables named tcs_rail_robotnum_index

where:

robotnum is the number of the robot being used.

index is the station number - 1.